## Slide 1: XCIST simulation & recon tests

## Slide 2: Baseline for CatSim phantom experiments
Scanner
1-mm x 1-mm uniform focal spot
1-mm x 1-mm detector pixels
1 slice
¼-detector column offset
Physics
120 kV
Quantum noise off
Electronic noise off
Scatter off
2x2 source sampling
2x2 detector sampling
2x view sampling
Protocol
300 mA
1-s rotation
360 views
Phantom
1024x1024x1-pixel phantom
0.5 mm per XY pixel
16 mm per Z pixel
Scaled to 0.5
Therefore, phantom is 256 mm XY and 8 mm Z.
Recon
1 slice
1-mm slice thickness
300-mm FOV
No offset
Standard kernel
HU

## Slide 3: Noise evaluation – 300 mA, spectrum-specific mu
No noise
Polyenergetic
Monoenergetic
Quantum noise
Quantum & electronic noise
Electronic noise
1 energy bin
Mean = 0.3 HU s.d. = 7.6 HU
Mean = 0.3 HU s.d. = 7.8 HU
Mean = 0.2 HU s.d. = 37.6 HU
Mean = 0.8 HU s.d. = 37.3 HU
Mean = 0.0 HU s.d. = 8.2 HU
Mean = 0.0 HU s.d. = 9.0 HU
Mean = 0.5 HU s.d. = 40.7 HU
Mean = 0.0 HU s.d. = 39.9 HU
Mean = 0.0 HU s.d. = 6.2 HU
Mean = 0.0 HU s.d. = 6.5 HU
Mean = -0.5 HU s.d. = 34.1 HU
Mean = 0.1 HU s.d. = 34.7 HU

## Slide 4: Noise evaluation – 100 mA, spectrum-specific mu
No noise
Polyenergetic
Monoenergetic
Quantum noise
Quantum & electronic noise
Electronic noise
1 energy bin
Mean = 0.0 HU s.d. = 7.8 HU
Mean = 0.1 HU s.d. = 13.1 HU
Mean = 0.9 HU s.d. = 66.4 HU
Mean = 0.3 HU s.d. = 65.9 HU
Mean = -0.1 HU s.d. = 6.5 HU
Mean = 0.1 HU s.d. = 20.2 HU
Mean = 1.4 HU s.d. = 69.7 HU
Mean = -1.4 HU s.d. = 70.1 HU
Mean = 0.0 HU s.d. = 5.3 HU
Mean = 0.3 HU s.d. = 12.8 HU
Mean = 0.3 HU s.d. = 60.7 HU
Mean = -0.2 HU s.d. = 62.4 HU

## Slide 5
No noise
Polyenergetic
Monoenergetic
Quantum noise
Quantum & electronic noise
Electronic noise
1 energy bin
Mean = 0.0 HU s.d. = 7.8 HU
Mean = 0.1 HU s.d. = 13.1 HU
Mean = 0.9 HU s.d. = 66.4 HU
Mean = 0.3 HU s.d. = 65.9 HU
Mean = -0.1 HU s.d. = 6.5 HU
Mean = 0.1 HU s.d. = 20.2 HU
Mean = 1.4 HU s.d. = 69.7 HU
Mean = -1.4 HU s.d. = 70.1 HU
Mean = 0.0 HU s.d. = 5.3 HU
Mean = 0.3 HU s.d. = 12.8 HU
Mean = 0.3 HU s.d. = 60.7 HU
Mean = -0.2 HU s.d. = 62.4 HU
Noise evaluation – 100 mA, spectrum-specific mu #2
Re-run no different
???

## Slide 6: Recon kernel evaluation
Soft
Bone
Standard
R-L
S-L
Only a slight difference between R-L and S-L.
Also, these shift the mean differently than the other 3 kernels (using the same mu).
Mean = 7.2 HU s.d. = 69.1 HU
Mean = 7.3 HU s.d. = 54.9 HU
Mean = 0.3 HU s.d. = 24.5 HU
Mean = -0.1 HU s.d. = 37.7 HU
Mean = -0.2 HU s.d. = 76.9 HU
ROI

## Slide 7: Recon kernel evaluation #2
Soft
Bone
Standard
R-L
S-L
Only a slight difference between R-L and S-L.
Also, these shift the mean differently than the other 3 kernels (using the same mu).
Mean = 7.2 HU s.d. = 69.1 HU
Mean = 7.2 HU s.d. = 56.1 HU
Mean = 0.2 HU s.d. = 24.1 HU
Mean = -0.2 HU s.d. = 38.7 HU
Mean = -0.3 HU s.d. = 78.5 HU
ROI
Re-run - these are different ???
They all use the same projections…

## Slide 8: Rotation time evaluation
0.5 s
2.0 s
1.0 s
Mean = 0.0 HU s.d. = 53.6 HU
Mean = 0.0 HU s.d. = 37.2 HU
Mean = 0.0 HU s.d. = 27.3 HU
ROI

## Slide 9: Rotation time evaluation #2
0.5 s
2.0 s
1.0 s
Mean = 0.0 HU s.d. = 53.6 HU
Mean = 0.0 HU s.d. = 37.2 HU
Mean = 0.0 HU s.d. = 27.3 HU
ROI
Re-run no different
???

## Slide 10: Number of views evaluation
100 views
1000 views
360 views
ROI 2
ROI 1
Mean = -9.8 HU s.d. = 133.5 HU
Mean = -9.6 HU s.d. = 44.1 HU
Mean = -9.8 HU s.d. = 32.0 HU
Mean = 0.6 HU s.d. = 39.0 HU
Mean = 0.2 HU s.d. = 38.6 HU
Mean = 0.9 HU s.d. = 65.5 HU

## Slide 11: Number of views evaluation #2
100 views
1000 views
360 views
ROI 2
ROI 1
Mean = -9.8 HU s.d. = 133.5 HU
Mean = -9.6 HU s.d. = 44.1 HU
Mean = -9.8 HU s.d. = 32.0 HU
Mean = 0.6 HU s.d. = 39.0 HU
Mean = 0.2 HU s.d. = 38.6 HU
Mean = 0.9 HU s.d. = 65.5 HU
Re-run no different
???

## Slide 12: Sampling evaluation
1 sample
Source
View
3 samples
2 samples
Detector
3 samples – 1 sample
No effect from detector sampling is expected

## Slide 13: Scatter evaluation
Scatter scale 1
No scatter
Scatter scale 8
Scatter scale 64

## Slide 14: Baseline
Center of eyes
16-row detector, 1-mm rows - > 9.095-mm Z FOV @ Z-axis
Center of eyes
outside of eye
Ca
Outside of eye
im
View 1 Detector col 0…900
View 91 Detector col 0…900
View 181 Detector col 0…900
View 271 Detector col 0…900
1024x1024x1-pixel phantom
Defined as 0.5 mm per XY pixel, and 16 mm per Z pixel, then scaled to 0.5. therefore, phantom is 256 mm XY and 8 mm Z.
CatSim but no eyes

## Slide 15: 16 slices, no phantom offset
16-slice recon, 1-mm slices, 300-mm FOV
We expect:
2 blank slices on each end (slices 1, 2, 15, 16)
3 slices on each half with phantom partially exposed; symmetrical (i.e., slices 3 = slice 14 ; slice 4 = slice 13 ; slice 5 = slice 12)
6 complete slices (slices 6-11)

## Slide 16: 16 slices, no phantom offset
16-slice recon, 1-mm slices, 300-mm FOV
300 mm
256 mm
Expected each end to have 2 blank + 3 partial, with partials symmetrical in Z.
Got 1 blank slice + 3 partial, with partials symmetrical in Z.

## Slide 17: 16 slices, phantom offset +50-mm X
View 1 Detector col 0…900
View 91 Detector col 0…900
View 181 Detector col 0…900
View 271 Detector col 0…900
1024x1024x1-pixel phantom
Defined as 0.5 mm per XY pixel, and 16 mm per Z pixel, then scaled to 0.5. therefore, phantom is 256 mm XY and 8 mm Z.
Center of eyes
16-row detector, 1-mm rows - > 9.095-mm Z FOV @ Z-axis
+X offset moves the phantom to the right – good!
Corner of eye
m

## Slide 18: 16 slices, phantom offset +50-mm X
300 mm
50 + 22 mm
So phantom offset is applied after the scaling – good!
Expected each end to have 2 blank + 4 partial, with partials symmetrical in Z.
Got 1 blank slice + 3 partial, with partials asymmetrical in Z.

## Slide 19: 16 slices, phantom offset +50-mm Y
View 1 Detector col 0…900
View 91 Detector col 0…900
View 181 Detector col 0…900
View 271 Detector col 0…900
1024x1024x1-pixel phantom
Defined as 0.5 mm per XY pixel, and 16 mm per Z pixel, then scaled to 0.5. therefore, phantom is 256 mm XY and 8 mm Z.
Center of eyes
16-row detector, 1-mm rows - > 9.095-mm Z FOV @ Z-axis
CatS
+Y offset moves the phantom up – good!
Bottom edge of CatSim

## Slide 20: 16 slices, phantom offset +50-mm Y
50 + 22 mm
16-slice recon, 1-mm slices, 300-mm FOV
Expected each end to have 2 blank + 4 partial, with partials symmetrical in Z.
Got 1 blank slice + 3 partial, with partials symmetrical in Z.
Slices 5-12

## Slide 21: 16 slices, phantom offset +4-mm Z
View 1 Detector col 0…900
View 91 Detector col 0…900
View 181 Detector col 0…900
View 271 Detector col 0…900
1024x1024x1-pixel phantom
Defined as 0.5 mm per XY pixel, and 16 mm per Z pixel, then scaled to 0.5. therefore, phantom is 256 mm XY and 8 mm Z.
Center of eyes
16-row detector, 1-mm rows - > 9.095-mm Z FOV @ Z-axis
Row 8 (and 1-7)
Row 9
+Z offset moves the phantom to a higher-numbered detector row – good!

## Slide 22: 16 slices, phantom offset +4-mm Z
9-16 all blank:
+Z offset moves the phantom to a lower-numbered slice – slice numbering needs to be reversed. Could do that by flipping the array, but more efficient if the array is written in the opposite Z order in C.
3-8 all complete:
Expected one end to have 2 blank + 4 partial + 2 complete,
Got 0 blank slices + 2 partial + 6 complete.

## Slide 23: 16 slices, phantom offset +8-mm Z
View 1 Detector col 0…900
View 91 Detector col 0…900
View 181 Detector col 0…900
View 271 Detector col 0…900
1024x1024x1-pixel phantom
Defined as 0.5 mm per XY pixel, and 16 mm per Z pixel, then scaled to 0.5. therefore, phantom is 256 mm XY and 8 mm Z.
Center of eyes
16-row detector, 1-mm rows - > 9.095-mm Z FOV @ Z-axis
Row 13
15
16
+Z offset moves the phantom to a higher-numbered detector row – good!
14

## Slide 24: 16 slices, phantom offset +8-mm Z
16-slice recon, 1-mm slices, 300-mm FOV
Expected one end to have 2 blank + 2 partial, with remaining 12 blank.
Got all blank slices except slice 1.

## Slide 25: 16 slices, no phantom offset
16-slice recon, 0.5-mm slices, 300-mm FOV
Expected each end to have 0 blank + 6 partial, with partials symmetrical in Z.
Got 0 blank slices + 3 partial, with partials symmetrical in Z.

## Slide 26: 16 slices, no phantom offset, +1-mm-Z recon offset
16-slice recon, 0.5-mm slices, 300-mm FOV
Looks almost OK but not quite. Slices 1-2 look OK (match slices 3-4 on prior slide) but slices 12-14 don’t match slices 14-16 on prior slide.

## Slide 27: Recon offset
No offset
Moved +22 mm X.
Spec’d +22-mm X
Moved 22 mm Y.
Spec’d +22-mm Y